BIPM GPS calibration information sheet

| Laboratory: |  |  | KRIS |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  | # 23 Mar. 2000(MJD 51616) 02:50 UTC |  |  |  |
| Date and hour of the end of measurements: |  |  | # 9 Apr. 2000(MJD 51643) 22:58 UTC |  |  |  |
| Receiver setup information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: NML |
|  Maker: | Allen-Osborne |  |  |  |  | Allen-Osborne |
|  Type: | TTR6 |  |  |  |  | TTR6 |
|  Serial number: | 415 |  |  |  |  | 467 |
|  Receiver internal delay (GPS) : | 50.0 |  |  |  |  |  |
|  Receiver internal delay (GLO) : |  |  |  |  |  |  |
|  Antenna cable identification: |  |  |  |  |  | NML IF |
| Corresponding cable delay : | 250.0 |  |  |  |  | 234.5 ns ± 0.5 ns |
|  UTC cable identification: |  |  |  |  |  |  |
| Corresponding cable delay : |  |  |  |  |  |  |
| Delay to local UTC : | 576 |  |  |  |  | 582 |
|  Receiver trigger level: |  |  |  |  |  |  |
|  Coordinates reference frame: | WGS-84 |  |  |  |  | WGS-84 |
| Latitude: | 36 23’ 18.0782” |  |  |  |  | 36 23’ 18.1” |
| Longitude: | 127 22’ 10.0564” |  |  |  |  | 127 22’ 10.058” |
| Height: | +123.94 m |  |  |  |  | +123.17 m |
| Antenna information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: |
|  Maker: | Allen Osborne |  |  |  |  | Allen Osborne |
|  Type: | TTR6 |  |  |  |  | TTR6 |
|  Serial number: |  |  |  |  |  | 572 |
| ## If the antenna is temperature stabilised |  |  |  |  |  |  |
|  Set temperature value : |  |  |  |  |  |  |
| Antenna cable information |  |  |  |  |  |  |
|  Maker: |  |  |  |  |  |  |
|  Type: |  |  |  |  |  |  |
|  Is it a phase stabilised cable: |  |  |  | No |  |  |
|  Length of cable outside the building : |  |  |  | 6.5 m |  |  |
| General information |  |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  |  | 16 ns |  |  |
| #  Is the laboratory air conditioned: |  |  |  | Yes |  |  |
|  Set temperature value and uncertainty : |  |  |  | 23 C  1 C |  |  |
|  Set humidity value and uncertainty : |  |  |  | 50 %  5 % |  |  |
| Cable delay control |  |  |  |  |  |  |
| Cable identification |  | delay measured by NML |  |  | delay measured by local method |  |
| NML-IF Antenna cable |  | 234.5 ns ± 0.5 ns |  |  |  |  |

Plot of the experiment set-up: ## Link to the local UTC of both receivers and Antenna positions

Description of the local method of cable delay measurement: