BIPM GPS calibration information sheet

| Laboratory: |  |  | SCL |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  | # 14/April/2000 (MJD51648) 0746 UTC |  |  |  |
| Date and hour of the end of measurements: |  |  | # 28/April/2000 (MJD51662) 0650 UTC |  |  |  |
| Receiver setup information |  |  |  |  |  |  |
|  |  | Local: SCL |  |  |  | Portable: NML |
|  Maker: |  | Allen Osborne |  |  |  | Allen Osborne |
|  Type: |  | TTR-6 |  |  |  | TTR-6 |
|  Serial number: |  | 417 |  |  |  | 267 |
|  Receiver internal delay (GPS) : |  | 55 ns |  |  |  | 68 ns |
|  Receiver internal delay (GLO) : |  | NA |  |  |  | NA |
|  Antenna cable identification: |  | SCL-IF Cable 6 |  |  |  | SCL-IF Cable 9 |
| Corresponding cable delay : |  | 728 ns |  |  |  | 720 ns |
|  UTC cable identification: |  | UTC(SCL) |  |  |  | UTC(SCL) |
| Corresponding cable delay : |  | 10 ns |  |  |  | 10 ns |
| Delay to local UTC : |  | NA |  |  |  | NA |
|  Receiver trigger level: |  | NA |  |  |  | NA |
|  Coordinates reference frame: |  | WGS84 |  |  |  | WGS84 |
| Latitude: |  | 22 ° 16 ' 47.5005 " |  |  |  | 22 ° 16 ' 47.4870 " |
| Longitude: |  | 114 ° 10 ' 22.7663 " |  |  |  | 114 ° 10 ' 22.7533 " |
| Height: |  | 186.86 metres |  |  |  | 186.86 metres |
| Antenna information |  |  |  |  |  |  |
|  |  | For Local Receiver: |  |  |  | For Portable Receiver: |
|  Maker: |  | Allen Osborne |  |  |  | Allen Osborne |
|  Type: |  | TTR-6 Antenna |  |  |  | TTR-6 Antenna |
|  Serial number: |  | 689 |  |  |  | 192 |
| ## If the antenna is temperature stabilised |  |  |  |  |  |  |
|  Set temperature value : |  | NA |  |  |  | NA |
| Antenna cable information |  |  |  |  |  |  |
|  Maker: |  | Suhner |  |  |  | Suhner |
|  Type: |  | RG213/U and RG58/U |  |  |  | RG214/U and RG58/U |
|  Is it a phase stabilised cable: |  | No |  |  |  | No |
|  Length of cable outside the building : |  | 41.8 metres |  |  |  | 41.8 metres |
| General information |  |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  |  | 5 ns |  |  |
| #  Is the laboratory air conditioned: |  |  |  | Yes |  |  |
|  Set temperature value and uncertainty : |  |  |  | (23 ± 1) °C |  |  |
|  Set humidity value and uncertainty : |  |  |  | (45 ± 8) % |  |  |
| Cable delay control |  |  |  |  |  |  |
| Cable identification | Delay measured by NML |  |  |  | Delay measured by local method |  |
| NML-IF Antenna cable | 234.5 ns ± 0.5 ns |  |  |  | 234.7 ns ± 0.5 ns |  |

Plot of the experiment set-up: Master Clock HP5071A Microphase Stepper Trak 6490 Distribution Amplifier HP5087A Time Scale Generator Trak 6460 NML Travelling GPS Receiver TTR-6 SCL Local GPS Receiver TTR-6 ## Link to the local UTC of both receivers and Antenna positions Antenna Antenna S/N: 192 S/N: 689 Roof Top (Outdoor) Antenna Antenna Cable 6 Cable 9 37F to 48F 36F 1 pps 1 pps UTC(SCL) UTC(SCL) 5 MHz 5 MHz 5 MHz 5 MHz 5 MHz Measurement Setup for the period: 14 April 2000 (MJD51648) to 28 April 2000 (MJD51662)

Description of the local method of cable delay measurement: ## Pulse Generator ## Time Interval Counter ## SR 620 Laboratory Frequency Standard System Start Stop P2 P3 50 Ohm Feedthrough Termination Tee Adapter Cable Under Test The port P2 was directly connected to the Start input of the Time Interval Counter. The cable under test was connected between the port P3 and the Stop input of the counter. The continuous square pulses generated from the Pulse Generator were splitted into two streams of square pulses by the Tee adapter. One stream was applied directly into the start input of the counter. The other stream was applied into the stop input of the counter through the cable under test. Both inputs of the counter were set in positive triggering at the mid-levels of the input pulses. The cable delay was measured by the time interval counter as the time delay between the start and stop inputs.