BIPM GPS calibration information sheet

| Laboratory: |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  |  |  |  |  |
| Date and hour of the end of measurements: |  |  |  |  |  |  |
| Receiver setup information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: NML |
|  Maker: | Allen-Osborne |  |  |  |  | Allen-Osborne |
|  Type: | TTR6 |  |  |  |  | TTR6 |
|  Serial number: | 461 |  |  |  |  | 467 |
|  Receiver internal delay (GPS) : | 50 ns |  |  |  |  |  |
|  Receiver internal delay (GLO) : |  |  |  |  |  |  |
|  Antenna cable identification: | TL IF |  |  |  |  | NML IF |
| Corresponding cable delay : | 229 ns |  |  |  |  | 234.5 ns ± 0.5 ns |
|  UTC cable identification: |  |  |  |  |  |  |
| Corresponding cable delay : | 6 ns |  |  |  |  | 6 ns |
| Delay to local UTC : | 45 |  |  |  |  | 45 ns |
|  Receiver trigger level: |  |  |  |  |  |  |
|  Coordinates reference frame: |  |  |  |  |  |  |
| Latitude: | 24 57 12.8957 |  |  |  |  | 24 57 12.8114 |
| Longitude: | 121 09 52.0651 |  |  |  |  | 121 09 52.0769 |
| Height: | 202.40 m |  |  |  |  | 202.18 m |
| Antenna information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: |
|  Maker: | Allen Osborne |  |  |  |  | Allen Osborne |
|  Type: | TTR6 |  |  |  |  | TTR6 |
|  Serial number: |  |  |  |  |  | 572 |
| ## If the antenna is temperature stabilised |  |  |  |  |  |  |
|  Set temperature value : |  |  |  |  |  |  |
| Antenna cable information |  |  |  |  |  |  |
|  Maker: |  |  |  |  |  |  |
|  Type: |  |  |  |  |  |  |
|  Is it a phase stabilised cable: |  |  |  |  |  |  |
|  Length of cable outside the building : |  |  |  | ~45 ft |  |  |
| General information |  |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  |  | < 4 ns |  |  |
| #  Is the laboratory air conditioned: |  |  |  | Yes |  |  |
|  Set temperature value and uncertainty : |  |  |  | 23 ± 5℃ |  |  |
|  Set humidity value and uncertainty : |  |  |  | no |  |  |
| Cable delay control |  |  |  |  |  |  |
| Cable identification |  | delay measured by NML |  |  | delay measured by local method |  |
| NML-IF Antenna cable |  | 234.5 ns ± 0.5 ns |  |  | 229 ns |  |

Plot of the experiment set-up: ## Link to the local UTC of both receivers and Antenna positions Antenna positions: TL: Lat: 24 57 12.8957 Lon: 121 09 52.0651 Hei: 202.40 m NML: Lat: 24 57 12.8114 Lon: 121 09 52.0769 Hei: 202.18 m

Description of the local method of cable delay measurement: HP 4195A configured as follows is used to measure the cable delay.