BIPM GPS calibration information sheet

| Laboratory: |  |  | Time and Frequency Laboratory - VMI |  |  |
| --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  | # 10-07-2001 |  |  |
| Date and hour of the end of measurements: |  |  | # 03-08-2001 |  |  |
| Receiver setup information |  |  |  |  |  |
|  | Local: |  |  |  | Portable: NML |
|  Maker: | Allen-Osborne |  |  |  | Allen-Osborne |
|  Type: | TTR6 |  |  |  | TTR6 |
|  Serial number: | 462 |  |  |  | 267 |
|  Receiver internal delay (GPS) : | 50 ns |  |  |  |  |
|  Receiver internal delay (GLO) : |  |  |  |  |  |
|  Antenna cable identification: | VMI IF |  |  |  | NML IF |
| Corresponding cable delay : | 250 ns |  |  |  | 234.5 ns ± 0.5 ns |
|  UTC cable identification: |  |  |  |  |  |
| Corresponding cable delay : |  |  |  |  | 7 ns |
| Delay to local UTC : | 38 |  |  |  | 61 ns |
|  Receiver trigger level: |  |  |  |  |  |
|  Coordinates reference frame: |  |  |  |  |  |
| Latitude: | 21 02 48.5460 |  |  |  | 21 02 48.2619 |
| Longitude: | 105 48 09.0600 |  |  |  | 105 48 09.6809 |
| Height: | -04.02 m |  |  |  | -05.94 m |
| Antenna information |  |  |  |  |  |
|  | Local: |  |  |  | Portable: |
|  Maker: | Allen Osborne |  |  |  | Allen Osborne |
|  Type: | TTR6 |  |  |  | TTR6 |
|  Serial number: |  |  |  |  |  |
| ## If the antenna is temperature stabilised |  |  |  |  |  |
|  Set temperature value : |  |  |  |  |  |
| Antenna cable information |  |  |  |  |  |
|  Maker: |  |  |  |  |  |
|  Type: |  |  | RG-58 |  |  |
|  Is it a phase stabilised cable: |  |  |  |  |  |
|  Length of cable outside the building : |  |  | ~130 ft |  |  |
| General information |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  | < 5 ns |  |  |
| #  Is the laboratory air conditioned: |  |  | Yes |  |  |
|  Set temperature value and uncertainty : |  |  | 23 ± 5 |  |  |
|  Set humidity value and uncertainty : |  |  | no |  |  |
| Cable delay control |  |  |  |  |  |
| Cable identification |  | delay measured by NML |  | delay measured by local method |  |
| NML-IF Antenna cable |  | 234.5 ns ± 0.5 ns |  |  |  |

Plot of the experiment set-up: ## PC ## PC 5 MHz 1 PPS TTR6(VMI) 5 MHz 1 PPS TTR6(NML) Frequency Distr. Amp 1 PPS Distr. Amp Frequency Converter VMI Equip. ## Link to the local UTC of both receivers and Antenna positions NFS UTC(VMI) D = 38 ns D = 61 ns * ** D = 7 ns Antenna positions: VMI: Lat: 21 02 48.5460 Lon: 105 48 09.0600 Hei: -04.02 m NML: Lat: 21 02 48.2619 Lon: 105 48 09.6809 Hei: -05.94 m D : Cable delay; * rise time of pulse < 5 ns; ** rise time of pulse < 20 ns.

Description of the local method of cable delay measurement: # U1-9 rise time of pulse < 1 ns # HP 53132 A Channel 1 Channel 2 R=75 HP 53132 A configured as follows is used to measure the cable delay. T-Connecor Cable