BIPM GPS calibration information sheet

| Laboratory: |  |  | NMI Sydney, Australia |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  | # 24/06/2004, MJD 53180 |  |  |  |
| Date and hour of the end of measurements: |  |  | # 22/07/2004, MJD 53208 |  |  |  |
| Receiver setup information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: NML |
|  Maker: | NMI/Topcon |  |  |  |  | NMI/Topcon |
|  Type: | Topcon Euro-80 |  |  |  |  | Euro-80 Dual Frequency |
|  Serial number: | 8RQRKXT534 |  |  |  |  | 8R633IOLON4 |
|  Receiver internal delay (GPS) : | 46.5 ns (uncalibrated) |  |  |  |  | 44.79 ns (uncalibrated) |
|  Receiver internal delay (GLO) : |  |  |  |  |  |  |
|  Antenna cable identification: | TCDF-1 |  |  |  |  | NML IF |
| Corresponding cable delay : | (75.9 ± 1.0) ns |  |  |  |  | (159.8 ± 1.0) ns |
|  UTC cable identification: | UTC(AUS) 9.1.02 |  |  |  |  | APMP Portable |
| Corresponding cable delay : | (76.0 ± 1.0) ns |  |  |  |  | (85.9 ± 1.0) ns |
| Delay to local UTC : | (76.0 ± 1.0) ns |  |  |  |  | (85.9 ± 1.0) ns |
|  Receiver trigger level: | 0.5V |  |  |  |  | 0.5 V |
|  Coordinates reference frame: | ITRF93 |  |  |  |  | ITRF2000 @ 27/06/04 |
| Latitude or X m | -4648200.298 |  |  |  |  | -4648199.675 |
| Longitude or Y m | 2560484.03 |  |  |  |  | 2560483.895 |
| Height or Z m | -3526505.358 |  |  |  |  | -3526506.097 |
| Antenna information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: |
|  Maker: | Topcon |  |  |  |  | Topcon/Javad |
|  Type: | Regant-1 |  |  |  |  | MarAnt |
|  Serial number: | RA0122 |  |  |  |  | MAGGD #0191 |
| ## If the antenna is temperature stabilised |  |  |  |  |  |  |
|  Set temperature value : |  |  |  |  |  | — |
| Antenna cable information |  |  |  |  |  |  |
|  Maker: |  |  |  | Rojone |  |  |
|  Type: |  |  |  | LMR400 |  |  |
|  Is it a phase stabilised cable: |  |  |  | No |  |  |
|  Length of cable outside the building : |  |  |  | 10m |  |  |
| General information |  |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  |  | ≤ 4ns |  |  |
| #  Is the laboratory air conditioned: |  |  |  | Yes |  |  |
|  Set temperature value and uncertainty : |  |  |  | (20 ± 2) °C |  |  |
|  Set humidity value and uncertainty : |  |  |  | (50 ± 10) % |  |  |
| Cable delay control |  |  |  |  |  |  |
| Cable identification |  | delay measured by NML |  |  | delay measured by local method |  |
| NML-IF Antenna cable |  | (159.8 ± 1.0) ns |  |  | (159.3 ± 1.0) ns |  |

Plot of the experiment set-up: ## Link to the local UTC of both receivers and Antenna positions

Description of the local method of cable delay measurement: We observe a typical day-to-day variation of ±0.5 ns in the delay measured for a given cable, and we therefore estimate the uncertainty of this method at ±1 ns.