BIPM GPS calibration information sheet

| Laboratory: |  |  | NMIJ |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  | # 2004-05-26 03:00 UTC |  |  |  |
| Date and hour of the end of measurements: |  |  | # 2004-05-31 04:00 UTC |  |  |  |
| Receiver setup information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: NML |
|  Maker: | AOA |  |  |  |  | NML/Topcon |
|  Type: | TTR6 |  |  |  |  | Euro-80 Dual Frequency |
|  Serial number: | 484 |  |  |  |  | 8R633IOLON4 |
|  Receiver internal delay (GPS) : | 50.0ns |  |  |  |  |  |
|  Receiver internal delay (GLO) : |  |  |  |  |  |  |
|  Antenna cable identification: |  |  |  |  |  | NML IF |
| Corresponding cable delay : | 259.0ns |  |  |  |  | (159.8 ± 1.0) ns |
|  UTC cable identification: |  |  |  |  |  |  |
| Corresponding cable delay : |  |  |  |  |  |  |
| Delay to local UTC : | 27.0ns |  |  |  |  | 510.6 ns |
|  Receiver trigger level: |  |  |  |  |  | 0.5 V |
|  Coordinates reference frame: | ITRF94 |  |  |  |  |  |
| Latitude: | 36 03 32.3826 (deg, min, sec) |  |  |  |  |  |
| Longitude: | 140 08 06.2173 (deg, min, sec) |  |  |  |  |  |
| Height: | 83.98 (m) |  |  |  |  |  |
| Antenna information |  |  |  |  |  |  |
|  | Local: |  |  |  |  | Portable: |
|  Maker: | AOA |  |  |  |  | Topcon/Javad |
|  Type: | GPS |  |  |  |  | MarAnt |
|  Serial number: | 682 |  |  |  |  | MAGGD #0191 |
| ## If the antenna is temperature stabilised |  |  |  |  |  |  |
|  Set temperature value : | — |  |  |  |  | — |
| Antenna cable information |  |  |  |  |  |  |
|  Maker: |  |  |  | Fujikura |  |  |
|  Type: |  |  |  | RG-55/U |  |  |
|  Is it a phase stabilised cable: |  |  |  |  |  |  |
|  Length of cable outside the building : |  |  |  | 15m |  |  |
| General information |  |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  |  | 3.7ns |  |  |
| #  Is the laboratory air conditioned: |  |  |  | Yes |  |  |
|  Set temperature value and uncertainty : |  |  |  | 23℃±1℃ |  |  |
|  Set humidity value and uncertainty : |  |  |  | 50％ |  |  |
| Cable delay control |  |  |  |  |  |  |
| Cable identification |  | delay measured by NML |  |  | delay measured by local method |  |
| NML-IF Antenna cable |  | (159.8 ± 1.0) ns |  |  |  |  |

TTR6 Carrier Distribution Amplifier (Epsilon AS) Pulse Distribution Amplifier (Agilent E8408A main frame with Carrier module E1752A ) Carrier Distribution Amplifier (HP5087A) NMIJ’s Master Clock (CC6) AOG (SigmaTau) model-110 *We are planning to introduce the AOG from this 3rd of June. E-80 Receiver (Calibration Trip) Carrier Distribution Amplifier (Agilent E8408A main frame with 1PPS module E1750A) E-80 Receiver (NMIJ) Carrier Distribution Amplifier (HP5087A) ASHTECH Z-12T Receiver Doubler 5MHz 1PPS 10MHz 10MHz 20MHz UTC(NMIJ) Plot of the experiment set-up: ## Link to the local UTC of both receivers and Antenna positions

Description of the local method of cable delay measurement: TIC Start Stop 1pps Cable