BIPM GPS calibration information sheet

| Laboratory: |  |  | TL |  |  |
| --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  | 2004/04/21 UTC 00:00 |  |  |
| Date and hour of the end of measurements: |  |  | 2004/05/02 UTC 23:59 |  |  |
| # Receiver setup information |  |  |  |  |  |
|  |  | Local: |  |  | ## Portable: NML |
|  Maker: |  | NML/Topcon |  |  | NML/Topcon |
|  Type: |  | Euro-80 Dual Frequency |  |  | Euro-80 Dual Frequency |
|  Serial number: |  |  |  |  | 8R633IOLON4 |
|  Receiver internal delay (GPS): |  | 45.1 ns |  |  |  |
|  Receiver internal delay (GLO): |  |  |  |  |  |
|  Antenna cable identification: |  |  |  |  | NML IF |
| Corresponding cable delay: |  | 119.1 ns |  |  | (159.8 ± 1.0) ns |
|  UTC cable identification: |  | RG-58 |  |  | RG-58 |
| Corresponding cable delay: |  | 30.7 ns |  |  | 37.6 ns |
| Delay to local UTC: |  | 0 ns |  |  | 0 ns |
|  Receiver trigger level: |  | 0.5 V |  |  | 0.5 V |
|  Coordinates reference frame: |  |  |  |  |  |
| Latitude: |  | 24.9535365344 |  |  |  |
| Longitude: |  | 121.1646005980 |  |  |  |
| Height: |  | 201.445 |  |  |  |
| # Antenna information |  |  |  |  |  |
|  | Local: |  |  | Portable: |  |
|  Maker: | Ashtech |  |  | Topcon/Javad |  |
|  Type: | Choke ring |  |  | MarAnt |  |
|  Serial number: | Ash701945C_M |  |  | MAGGD #0191 |  |
| If the antenna is temperature stabilized |  |  |  |  |  |
|  Set temperature value: |  |  |  |  |  |
| # Antenna cable information |  |  |  |  |  |
|  Maker |  |  | Beldon |  |  |
|  Type |  |  | RG-8 |  |  |
|  Is it a phase stabilised cable: |  |  | No |  |  |
|  Length of cable outside the building: |  |  | 15 m |  |  |
| # General information |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  |  |  |  |
|  Is the laboratory air conditioned: |  |  | Yes |  |  |
|  Set temperature value and uncertainty: |  |  | 23 ± 1 C |  |  |
|  Set humidity value and uncertainty: |  |  | 50 ± 5 % |  |  |
| # Cable delay control |  |  |  |  |  |
| Cable identification | delay measured by NML |  |  | delay measured by local method |  |
| NML-IF Antenna cable | (159.8 ± 1.0) ns |  |  | (159.0 ± 1.5) ns |  |

BIPM GPS calibration information sheet

| Laboratory: |  |  | TL |  |  |
| --- | --- | --- | --- | --- | --- |
| Date and hour of the beginning of measurements: |  |  | 2004/04/21 UTC 00:00 |  |  |
| Date and hour of the end of measurements: |  |  | 2004/05/02 UTC 23:59 |  |  |
| # Receiver setup information |  |  |  |  |  |
|  |  | Local: |  |  | ## Portable: NML |
|  Maker: |  | AOA |  |  | NML/Topcon |
|  Type: |  | TTR6 |  |  | Euro-80 Dual Frequency |
|  Serial number: |  |  |  |  | 8R633IOLON4 |
|  Receiver internal delay (GPS): |  | 50 ns |  |  |  |
|  Receiver internal delay (GLO): |  |  |  |  |  |
|  Antenna cable identification: |  | RG-58 |  |  | NML IF |
| Corresponding cable delay: |  | 308 ns |  |  | (159.8 ± 1.0) ns |
|  UTC cable identification: |  | RG-58 |  |  | RG-58 |
| Corresponding cable delay: |  | 5 ns |  |  | 37.6 ns |
| Delay to local UTC: |  | -46.2 ns |  |  | 0 ns |
|  Receiver trigger level: |  | 0.5 V |  |  | 0.5 V |
|  Coordinates reference frame: |  |  |  |  |  |
| Latitude: |  | 24.9535356144 |  |  |  |
| Longitude: |  | 121.1645396555 |  |  |  |
| Height: |  | 201.631 |  |  |  |
| # Antenna information |  |  |  |  |  |
|  | Local: |  |  | Portable: |  |
|  Maker: | AOA |  |  | Topcon/Javad |  |
|  Type: |  |  |  | MarAnt |  |
|  Serial number: |  |  |  | MAGGD #0191 |  |
| If the antenna is temperature stabilized |  |  |  |  |  |
|  Set temperature value: |  |  |  |  |  |
| # Antenna cable information |  |  |  |  |  |
|  Maker |  |  |  |  |  |
|  Type |  |  | RG-58 |  |  |
|  Is it a phase stabilised cable: |  |  | No |  |  |
|  Length of cable outside the building: |  |  | 15 m |  |  |
| # General information |  |  |  |  |  |
|  Rise time of the local UTC pulse: |  |  |  |  |  |
|  Is the laboratory air conditioned: |  |  | Yes |  |  |
|  Set temperature value and uncertainty: |  |  | 23 ± 1 C |  |  |
|  Set humidity value and uncertainty: |  |  | 50 ± 5 % |  |  |
| # Cable delay control |  |  |  |  |  |
| Cable identification | delay measured by NML |  |  | delay measured by local method |  |
| NML-IF Antenna cable | (159.8 ± 1.0) ns |  |  | (159.0 ± 1.5) ns |  |

Plot of the experiment set-up: Link to the local UTC of both receivers and Antenna positions Lat: 24.9535356 Lon: 121.1645397 Hgt: 201.631 Lat: 24.9535365 Lon: 121.1646006 Hgt: 201.445 UTC (TL) 1PPS Distribution AOA TTR6 Portable NML/Topcon Local NML/Topcon -41.2 ns 30.7 ns 37.6 ns 308 ns 119.1 ns 159.8 ns

Description of the local method of cable delay measurement: Using HP network analyzer for measuring cable delay.